## Slide 1: Yangtian Yan (System Application Engineer)
High-Precision Temperature Measurement System Based on ADUCM360 Integrated with Temperature Compensation Algorithm

## Slide 2
Background / RTD
Resistance Temperature Detector (RTD) Overview
Most stable, most accurate, rugged easy to connect and implement. Great temperature range -240℃ ~ 650℃.
Figure 1 . Pt100 RTD Resistance from -200℃ to 800℃
  For positive temperatures:
  For negative temperatures from -200℃ ~ 0℃:
  The coefficients in the CVD equations are defined by IEC-60751 standard. is the resistance of RTD at 0℃. For a European standard, the coefficients are as follows:
  ΔT=ΔR/0.385
Approximate Expression
T represents the measuring temperature
ΔT represents the change of measuring temperature value

## Slide 3
System configuration / Parameter selection
RTD 4-Wire Configuration
The ratiometric measurement method
feature | Configuration
External Reference | Enable
VREF+ | Selected as buffer inputs
VREF- | Selected as buffer inputs
Headroom Resistor | 150Ω(0.1% 10PPM/℃)
Reference Resistor | 2.5kΩ(0.1% 10PPM/℃)
ADC Mode | Timer trigger single ADC conversion mode /ADC DMA mode
feature | Configuration
Gain | 4
Filter | Sin3
Current | 1mA/50uA
ADUCM360
1. “ How to Select and Design the Best RTD Temperature Sensing System ”
2. “ Thermistor-Based Temperature Sensing System—Part 1: Design Challenges and Circuit Configuration ”

## Slide 4
Calibration Measurement / Comparison of different Solutions
Test Result (1mA)
Comparison of Test result (1mA) and measurement result of CN0221 (200uA)
Test Equipment
FLUKE 7625A
The schematic and test result of Demo CN0221

## Slide 5
RTD Measurement / Self-heating Calculation
Using 1mA/50uA Excitation Current with Class AA RTD
to measure the water temperature
By designing a multi-physics coupling model involving fluid dynamics, fluid-solid heat transfer, and electrothermal processes, we can compute and predict the temperature rise of the RTD resistor.“
The temperature/heat transfer field governing equation
Current value | Configuration | Reference voltage
1mA | Gain=4，reference resistor=2.5kΩ, | 2.5V
50uA | Gain=128, reference resistor=50kΩ | 2.5V
Current value | Maximum measurement error | Self-heating error
1mA | 0.0799℃ | 0.051℃
50uA | 0.196℃ | 0.015℃
Measurement error analysis
R
The fluid/laminar field governing equation
The electric current field governing equation
v
Fig.1
Fig.2

## Slide 6
Simulation model Design and Accuracy Verification
Fig.1
Mode Design
Fig.8
Fig.2
Fig.3
Fig.4 Continuous mode (1mA)
Fig.5 Discontinuous mode (current=1mA) sampling period=20.6s
Current source turn-on duration=600ms
Current source turn-off duration= 20s
Fig. 6
Continuous mode test/simulation comparison
Fig.7
Fig.9
Fig.10
Fig.11

## Slide 7
Final test results using the compensation algorithm
Six measurement mode with self-heating prediction curve
Fig.1
Fig.2 Measured results of different current duty cycles
By using the compensation algorithm.
Fig.3 Discontinuous mode 1 result
Fig.4 Discontinuous mode 2 result
Fig.5 Discontinuous mode 3 result
Fig.6 Discontinuous mode 4 result
In all 5 modes, the maximum error is less than 0.03℃